## Slide 1: Validation of 2 chest X-Ray AI algorithms on Singular dataset CARPL.ai
Presented by – Pipladh Arora
Mentor – Dr. Siddharth Shekhar Mishra
IIMHR, Delhi

## Slide 2: Introduction
What is AI?
AI in medical imaging
Performance testing of AI
Requirement for validation
Testing on a new Dataset
Why on new dataset?

## Slide 3: Objectives
To compare performance of two algorithms
To test the generalizability of two algorithms
To identify the areas of improvement in the algorithm
To find the optimum threshold value for each algorithm
To conduct efficiency testing of the algorithms

## Slide 4: Methodology
The study was carried out at CARPL.ai.Inc. It is a descriptive study design with secondary research.
This study includes 2 AI models compared on same dataset which were performed from 7th March 7th June using CARPL platform as a tool for validating various AI models on different types of medical images.
Research articles were extracted using PubMed, google scholar and various other databases that included research studies on validation of AI models that investigate diagnostic decisions/prediction of medical images.
Inclusion Criteria:
Research studies on validation of AI models
Challenges faced during performance testing of AI models
Computing ways to improve AI algorithms
Exclusion Criteria: Articles that did not involve evaluation of performance of AI algorithms on medical images.

## Slide 5
Results (Algorithm 1)
AUC for Algorithm 1 is 0.894, and as explained closer the AUC to 1, better is the performance of the model. Here the model gives the probability of approx. 89.4% that it can detect images with different abnormalities correctly.

## Slide 6
Results (Algorithm 1)
The above images depict that in case of Algorithm 1, there are no Far North cases and the AI inference is in compliance with the Ground Truth.
Results (Algorithm 1)

## Slide 7
Results (Algorithm 1)
The above are the Far South cases, i.e the AI probability is way to inaccurate in detecting the abnormality and marked a positive case as highly negative. Again, these are also the cases that the AI is lagging at and we send such cases back to the AI Developers so that they can make the appropriate changes and rectify the errors where the AI is lagging in order to make the algorithm more generalized and way more accurate.

## Slide 8
Results (Algorithm 2)
This model gives an AUC value of 0.79 for Abnormality Detection. This means the model can detect abnormal chest X-Rays with a probability of 79%.
This algorithm best performs for this abnormality at a threshold of 40, giving F1 score of 0.79 and specificity of 0.73 and sensitivity of 0.71

## Slide 9
Results (Algorithm 2)
The above image shows the Far North cases for Algorithm 2. As it is clearly visible that in case of algorithm 1 the GT and the AI inference match but the inference result of Algorithm 2 is highly inaccurate. These are the cases that are then sent back to the AI development team so that the errors can be worked upon and the AI can be made more robust and accurate.

## Slide 10
Results (Algorithm 2)
The above are the Far South cases, i.e the AI probability is way to inaccurate in detecting the abnormality and marked a positive case as highly negative. Again these are also the cases that the AI is lagging at and we send such cases back to the AI Developers so that they can make the appropriate changes and rectify the errors where the AI is lagging in order to make the algorithm more generalized and way more accurate.

## Slide 11
Upon comparing the AUC values of the two models only for Abnormality Detection, then it clearly states that AI model 1 performs better and has greater probability of detecting the abnormality correctly.
Comparison

## Slide 12
Result
Validation Imperatives The two fundamental factors for conducting a validation are Datasets and clinical proficiency. To evaluate an AI algorithm an extensive amount of dataset is required which are clinically significant in order to outstand as how the algorithm performs in different situations. In addition to large datasets, clinical judgment is also an important factor as it helps to understand about the performance of an algorithm. In many situations it is difficult to evaluate why an algorithm failed. In such instances help of a clinical expertise is compelling on analyzing failed cases and provide reasons for it.

## Slide 13
Absolute Validation
For any AI algorithm to be used in a clinical domain, its generalizability is critical and also a challenge. During the development of an algorithm a dataset is used which is known as training dataset that is used for fine-tuning the parameters of the algorithm. Addition to training dataset, a testing/validation dataset is required for evaluating the performance of the algorithm. If an external dataset is used for validation it is known as external validation of the algorithm. Since the model would perform better on internal dataset, an external validation is critical for determining the performance the algorithm before deploying it in a clinical practice. If an algorithm is developed, trained and tested from using the data of one site it may not provide the accurate results of the performance of the AI model.
Result

## Slide 14
Selection of the Data Mix
Once the ground truth is prepared it is critical to evaluate the mix of cases to validate the AI model. Generally, AI model provides two such type of failures, one being False Positives, where the model gave an output of the image having the abnormality but in reality, it was negative, second is False Negative where AI gave an output for the image being normal instead it was abnormal in reality. Now to study the FP rate, a dataset without many positives are required as to evaluate how often did the model missed negative cases and stated them as positives. Similarly, to study the rate of FN, a dataset without many negative cases are required, to determine how frequently the algorithm did miss actual positives cases and stated them as negatives.
Result

## Slide 15
Limitations
The generalizability and the validation of Medical Imaging AI algorithms, like the ones that are being tested upon on this study depend on the volume of the data that it runs through. Thus, the results of this study may have a intrinsic bias due to the dataset that was used.
The compliance of the algorithms vary among themselves, as a result, certain findings may be better expressed in one over the other. This however is innate to the AI algorithm itself.
The Ground Truth which is used as the reference to validate the AI findings was collected and verified by a Senior Radiologist. This may have brought in some unintentional human based error or lapse in judgement.

## Slide 16
conclusion
In the world of AI, the term validation refers to fine-tuning the parameters for algorithm development and the test is used to evaluate the algorithm performance.
There are two aspects to evaluate the performance of any predictive model i.e Discrimination and Calibration which are usually determined by plotting ROC curve and Calibration curve respectively.
Another critical or one of the essentials for any validation of an AI model is external validation which increases the robustness of the algorithm. An extensive amount of data should be considered for evaluation of any model.
Lastly, how the model would benefit patients in long run should be a point of concern, which can be avoided by performing number of clinical trials on the algorithm before deploying it on a clinical practice.

## Slide 17
Kim DW, Jang HY, Kim KW, Shin Y, Park SH. Design characteristics of studies reporting the performance of artificial intelligence algorithms for diagnostic analysis of medical images: Results from recently published papers. Korean J Radiol [Internet]. 2019 [cited 2022 Jun 20];20(3):405–10. Available from: http://dx.doi.org/10.3348/kjr.2019.0025
Park SH, Han K. Methodologic guide for evaluating clinical performance and effect of artificial intelligence technology for medical diagnosis and prediction. Radiology [Internet]. 2018;286(3):800–9. Available from: http://dx.doi.org/10.1148/radiol.2017171920
Mahajan V, Venugopal V, Gaur S, Gupta S, Murugavel M, Mahajan H. The Algorithmic Audit: Working with vendors to validate radiology-AI algorithms -how we do it [Internet]. Vixra.org. [cited 2022 Jun 20]. Available from: https://vixra.org/pdf/1909.0104v1.pdf
Hand DJ, Khan S. Validating and verifying AI systems. Patterns (N Y) [Internet]. 2020;1(3):100037. Available from: https://www.sciencedirect.com/science/article/pii/S2666389920300428
Tao C, Gao J, Wang T. Testing and quality validation for AI software–perspectives, issues, and practices. IEEE Access [Internet]. 2019;7:120164–75. Available from: http://dx.doi.org/10.1109/access.2019.2937107
references

## Slide 18
Thank you

## Slide 19
Skills/topics that I’ve learned
Topics
  AI
  Deep Learning
  AI in medical imaging
  HIPAA
  Various Legal Agreements.
Skills
  Amazon Quicksight
  PyCharm
  Meddiff PACS
  Innvowave PACS